## Slide 1
Acid distillation and solvent extraction methods1,2,3,4 are widely used to determine monomethyl mercury (MMHg) content in sediments. During initial phases of the Penobscot River Mercury Study in Maine, USA in 2006 – 07, both methods were evaluated for performance. A combination of results from 3 laboratories (Flett Research Ltd., Battelle Marine Sciences Laboratory and Trent University) indicated the following:
  Results from solvent extraction average ~50% lower than acid distillation (number of tested samples >200);
  CRM IAEA-405 and matrix spike recoveries were similar and good with both methods;
  Artifact generation from conversion of Hg2+ to MMHg in acid distillation2,4,5 is low (<0.1%) according to isotopic labeling method6 and non-isotopic standard addition Hg2+ spike recovery and therefore is not a major factor for higher distillation results.
The cause of the discrepancy was unclear and therefore it was decided to test the following hypothesis: Distillation is providing the correct methyl mercury concentrations and the application of heat and/or acid is required to release methyl mercury for accurate determination by the solvent extraction method.
Two large well-homogenized sediment samples (one river and one estuary, taken as part of QA/QC procedures for the Penobscot Study) were used to test this hypothesis.
1. M. Horvat, Water Air Soil Pollut. 56 (1991) 95; 2. M. Horvat, N.S. Bloom, L. Liang, Anal. Chim. Acta 281 (1993) 135; 3. M. Horvat, N.S. Bloom, L. Liang, Anal. Chim. Acta 282 (1993) 153; 4. Bloom, N.S., Coleman, J.A., and Barber, L., Fres. Anal. Chem. 358 (1997) 371; 5. Karl C. Bowles, Simon C. Apte, Anal. Chim. Acta 419 (2000) 145; 6. H. Hintelmann, Chemosphere 39 (1999) 1093
MMHg determined by KBr Solvent Extraction underestimates the real MMHg concentration in Penobscot River sediment samples by ~50% due to low extraction efficiency. A portion of MMHg in these sediments is in a form that is different than the MMHg in standard spike solutions, making it unavailable to solvent extraction at room temperature: matrix spike recovery is good but final measurement is low. Good spike recovery does not guarantee good sample recovery.
Pretreatment of sediment with heat, and particularly heat in the presence of KCl/H2SO4 ,significantly improves the solvent extraction process. Although the conversion of inorganic Hg2+ is correspondingly increased, from 0.03% to 0.13%, the artifact accounts for less than 7% of the overall MMHg detected, and is only significant when THg/MeHg ratio is extremely high.
Poor solvent extraction is matrix specific. Solvent extraction of CRM IAEA-405 is nearly complete while test sediments ES2 and OB2 extractions were only about 50%. Good CRM recovery does not guarantee good sample recovery.
Comparison of Two Methods for the Measurement of Methyl Mercury Concentrations in Penobscot River Sediments
Dr. R.J. Flett, Flett Research Ltd; Ms. Brenda Lasorsa, Battelle Marine Sciences Laboratory;
Dr. Gary Gill, Battelle Marine Sciences Laboratory; Dr. Holger Hintelmann, Trent University
Introduction
Methods
Results
Results (Continued)
3. Solvent extraction MMHg results on ES2 and OB2 wet sediment are about half of those determined by distillation even though Matrix Spike recoveries and CRM IAEA405 recoveries were all better than 90% with either method.
6. Heat pretreatment of sediments at 100oC prior to solvent extraction
6.1 Heat ES2 wet sediment with KCl/H2SO4 OR in DI only
Conclusions
Acknowledgment
References
Sediment
QA/QC
Hg2+ spike
1. DISTILLATION with KCl/H2SO4
147oC distillation,
Collect 80~90% distillate
0.1% NaBEt4
derivatization
Purge & Trap
GC-CVAFS
2. REGULAR SOLVENT EXTRACTION with KBr/H2SO4/CH2Cl2
+ CH2Cl2 shake 1hr
Centrifuge 10min
0.1% NaBEt4
derivatization
Purge & Trap
GC-CAVFS
2ml CH2Cl2 layer, 15ml DI
45-50oC evaporation
N2 purging
3.1 Distill the previously extracted sediment
3.2 Extract the previously extracted sediment
3.3 Sample pretreatment with 100oC heat and DI OR 100oC heat and KCl/H2SO4 OR Sonication
+ CH2Cl2 shake 1hr
Centrifuge 10min
The remaining steps of extraction as in Section 2
Extracted sediment purged of trace CH2Cl2
Go through Acid Distillation procedure
Sonication, or
Heat at 100oC with DI, or
Heat at 100oC with KCl/H2SO4
4. INORGANIC Hg2+ SPIKE CONVERSION TO MMHg
1. Some of the Penobscot River Mercury Study data that prompted the present investigation. All results corrected for spike recovery. Acid distillation recovered MMHg at an average of 2.13 times of that recovered by solvent extraction.
4. Sonication pretreatment of wet ES2 sample before going through solvent extraction didn’t improve MMHg recovery. The MMHg detected was still half of that determined by distillation.
5. After a solvent extraction of wet ES2, the previously extracted sediment was distilled or went through a 2nd solvent extraction.
Sample Type | MMHg detected in Sediment by Distillation, ng/gdw | MMHg detected in sediment by KBr extraction, ng/gdw | MMHg detected in previously KBr solvent extracted sediment (converted to equivalent MMHg concentration in dry sediment) | 
 |  |  | Distillation, ng/gdw | KBr Extraction, ng/gdw
Wet ES2 | 28.1 | 12.4 | 7.61 | 2.23
Wet ES2 with MMHg spike |  | 42.55 | 7.35 | 4.13
IAEA405 (5.49±0.53 ng/g) | 5.41 | 4.48 | 0.31 | 0.40
Hg2+ conversion to MMHg | 0.07% | 0.03% | 0.08~0.16% | <0.01%
Sediments from two sites in Penobscot River, Maine, USA were sampled in this study. ES2 was obtained from the river estuary, whereas OB2 was taken upstream from the estuary. In the flow chart below, DI is ion-exchanged pure water, KCl/H2SO4 is 1.6% KCl in 8 M H2SO4 solution, CuSO4 is 1M CuSO4 solution, KBr/H2SO4 is 18% KBr in 10% H2SO4 solution, CVAFS is cold vapor atomic fluorescence spectrometry, QA/QC includes MMHg matrix spike, inorganic Hg2+ spike and sediment CRM IAEA405. When wet sediments were analysed, all MMHg concentration was normalized into dry weight basis, as ng/gdw.
6.2 Effect of different heating times on ES2 and OB2 wet sediment with KCl/H2SO4
Wet ES2 and OB2 showed similar trends of heating effect.
Comparative results from Distillation are 28.1ng/gdw and 21.2 ng/gdw for ES2 and OB2 wet, respectively.
The matrix spike and CRM recoveries were good.
MMHg solvent extraction recovery was highest when sediment was pretreated with heat @ 100oC with KCl/H2SO4.
If preheated in DI only, MMHg detected were higher than regular KBr/CH2Cl2 extraction, but still significantly lower than that from distillation.
The MMHg conversion from Hg2+ with this heating process was 0.07%.
The Penobscot River Mercury Study is thanked for permitting these data to be presented here.
Chart: Solvent Extraction - MMHg (ng/gdw)
Second KBr extraction as in Section 2 above
+ KCl/H2SO4
Sediment
QA/QC
Hg2+ spike
+ CuSO4
+ KBr/H2SO4
Sediment
QA/QC
Hg2+ spike
+ CuSO4
+ KBr/H2SO4
2. Trace level isotopic Hg labeling experiments showed 0.01% to 0.04% inorganic Hg2+ converted into MMHg while high level non-isotopic Hg2+ spike experiments showed 0.07% conversion occurred in the distillation of Penobscot River sediments, vs 0.03% found in extraction. The artifact MMHg determined by isotopic labeling procedure was less than 2.4% of the MMHg value determined by the distillation method.
The distillation of previously extracted sediment detected 7.61 ng/gdw MMHg, indicating a significant amount of MMHg remained after extraction. On the other hand, the distillation of previously extracted sediment which contained a matrix spike found a similar amount of MMHg (7.35) as in the unspiked sample. This indicated the aqueous based MMHg in matrix spike is easily extracted into CH2Cl2 (>90% recovery was found) whereas ambient MMHg in ES2 is poorly extracted.
Sample Station | Sample Description | Ambient THg*, ng/gdw | Ambient MMHg, ng/gdw | Average 200Hg methylated (artifact), %
OV5 (n=3) | freshwater | 51 | 0.72 | 0.01
OB2 (n=3) | tidal river | 1022 | 25.0 | 0.02
OB4 (n=6) | tidal river | 1078 | 15.1 | 0.03
ES2 (n=6) | estuary | 1622 | 27.3 | 0.04
*: THg results listed here are site average of 2007 from all labs
~0.4g ES2 wet sediment, heat with 0.2ml KCl/H2SO4 in 4ml DI or without KCl in 1 or 4 ml DI before extraction.
~0.4g ES2 and OB2 wet sediment, heat with 0.2ml KCl/H2SO4 in 1ml DI (not 4ml) before extraction.
3. MODIFICATIONS of the REGULAR SOLVENT EXTRACTION
For each experiment above, External Hg2+ spike was included to evaluate the MMHg conversion. Hg2+ spike is more than 1000 X MMHg concentration measured from Acid Distillation.
ES2 and OB2 sediment samples used in this study have Total Hg of 892 and 1185 ng/gdw respectively, MMHg (by distillation) of 28.1 and 21.2 ng/gdw respectively. With typical 0.07% conversion found in distillation, the artifact MMHg generated (as a % of ambient) is 2.2% for ES2 and 3.9% for OB2. This artifact is not a major contributor to the large discrepancy between the two methods.
Chart: MMHg concentration, ng/gdw
| Solvent extraction with different sample pretreatment
Heat 30', KCl,1ml DI | 24.1
Heat 90', 1ml DI only | 17.93
Heat 30', 4ml DI only | 13.4
 | 0
IAEA405, heat 90', 1ml DI only | 5.29
 | 0
Poster presented at the 10th International Conference on Mercury as a Global Pollutant (ICMGP), Halifax, Nova Scotia, Canada, July 24-29, 2011
Chart: MeHg concentration, ng/gdw
| Analysis with wet ES2 Analysis with wet OB2
Wet ES2 sediment, Acid Distillation | 28.1
Wet ES2 Sediment,KBr Extraction | 11.65
IAEA405, Acid Distillation | 5.41
IAEA405, KBr Extraction | 5.46
 | 0
Wet OB2 sediment, Acid distillation | 21.15
Wet OB2 sediment, KBr extraction | 10.5
IAEA405 distillation along with OB2 | 5.23
IAEA405 extracted along with OB2 | 5.15
 | 0
IAEA405 book value, 5.49±0.53ng/g
Spike recovery
Spike recovery
CRM recovery
CRM recovery
Chart: Matrix spike and CRM recovery, %
| Analysis with wet ES2 Analysis with wet OB2
Wet ES2 sediment, Acid Distillation | 95.9
Wet ES2 Sediment,KBr Extraction | 96
IAEA405, Acid Distillation | 98.5
IAEA405, KBr Extraction | 99.5
 | 0
Wet OB2 sediment, Acid distillation | 93.6
Wet OB2 sediment, KBr extraction | 90.2
IAEA405, Acid distillation | 95.3
IAEA405, KBr Extraction | 93.8
Results from May and July 2007 Sediments, all core depths included